## Slide 1: CUTE – A low background facility for testing cryogenic detectors Installation Review Phase 2A
Serge Nagorny, CUTE Project Manager
on behalf Queen’s team of the CUTE facility
September 19th, 2018

## Slide 2: Motivation
Test and validation of entire SuperCDMS detector tower (6 detector + electronics)
Measurements in a low background environment avoiding cosmogenic activation of detector’s material (3H, 32Si, etc.)
Complete detector characterization to understand its intrinsic background and noise issues
Confirmation of screening program and handling procedures
Early science run for dark matter search can be performed thanks to low-background environment and low-energy threshold of SuperCDMS detectors

## Slide 3: General view of the CUTE facility
Monorail with
1000 kg hoist
Platform
Movable
HDPE shield
UP water
shield tank
Cryostat
Cleanroom
Drywell with
Pb shield/
Mu-metal

## Slide 4: Arrangement
3/9/2016
Recent changes in the CUTE facility layout

## Slide 5: Installation Review Phase 2A will include following items:
PE shield installation
Pb shield installation (including mu-metal)
Clean room installation
Network
All actions planned to be completed by Dec 2018

## Slide 6: PE shielding
20 cm thick HDPE layer on top of the cryostat for neutron flux suppression
Produced, currently stored at site

## Slide 7: PE shield
8 layer of 1” HDPE sheet
Rigging/lifting test at producer site

## Slide 8: Lead shielding Produced, commissioning - Oct 2018, delivery to SNOLAB - end Oct 2018/beginning Nov 2018
Pb inner layer
< 3 Bq/kg
Pb outer layer
< 30 Bq/kg
650 kg
140 kg
2 cm
inner layer
9 cm
outer layer
13 cm
outer layer

## Slide 9: Cleanroom
Space for two operators
Aluminum structure with LEXAN panels
Resistant to seismic events
Movable
Lifting device for cryostat cans
Low Rn air supply (< 3 Bq/m3)
Later SuperCDMS Rn-filter system (< 0.1 Bq/m3) will be used
Crane moves cryostat between cleanroom and drywell

## Slide 10: Cleanroom
Bottom steel frame
Examples of Aluminum anodized frames

## Slide 11: Electrical/Network diagram
CUTE workstation
Electrical panel board
DAQ rack
Cleanroom/top
Cleanroom/inside
Platform/top
power
phone
ethernet
power
phone
ethernet
phone
ethernet
phone
ethernet
power
power
plug

## Slide 12
Start '18 Jun 01
Finish '19 Apr 01
'18 Jun
'18 Jul
'18 Aug
'18 Sep
'18 Oct
'18 Nov
'18 Dec
'19 Jan
'19 Feb
'19 Mar
6.6 Electrical Power Distribution '17 Nov 30 - '18 Sep 17
6.16 Clean Room Fab/Install '18 Aug 08 - '18 Oct 26
6.14 Cryostat Design/Installation '18 Jul 24 - '19 Mar 26
6.11 Cryostat & Equipment Shipping '18 Nov 05 - '18 Dec 14
6.10 Network and Communications '18 Aug 08 - '18 Oct 15
6.9 Instrument Air '18 Jul 24 - '18 Nov 12
6.17 PE Shield install '18 Aug 08 - '18 Oct 12
6.18 Magnetic Shield Design/Install '18 Aug 08 - '18 Dec 20
6.19 Lead Shield Install '18 Aug 08 - '18 Dec 20
6.20 Gamma System '18 Aug 31 - '19 Feb 28
6.23 Temp Water Chiller/Compressor '18 Oct 17 - '18 Nov 27
Cryostat Installation '19 Feb 13 - '19 Mar 26
6.2 Liners and Lid Installation '18 Sep 24 - '18 Oct 10
6.15 Instrumentation Rack Installation '19 Jan 30 - '19 Feb 05
6.12 Dilution GHS Installation '19 Jan 30 - '19 Feb 12
6.13 PT Compressor install '19 Jan 30 - '19 Feb 12
6.22 DAQ computer install '19 Jan 30 - '19 Feb 05
SNOLAB Approval - Phase 2A Installation '18 Oct 10
Clean Room at SNOLAB '18 Sep 24
Clean Room Final Inspection '18 Oct 26
Cryostat & Equip. at SNOLAB '18 Nov 30
Installation Review - Phase 2B Infrastructure '19 Jan 15
SNOLAB Approval - Phase 2B Installation '19 Jan 29
Operation Readiness Review '19 Mar 12
SNOLAB Approval Operations '19 Mar 26
PE Shield at SNOLAB '18 Aug 08
Lead Shield at SNOLAB '18 Nov 12
Magnetic Shield at SNOLAB '18 Sep 18
Tank Fill Complete '18 Oct 10
Suspension System at SNOLAB '18 Dec 21
Cryostat inst

## Slide 13: CUTE tasks till Operation with SuperCDMS detectors (1)
Step | Time required
Suspension system test @ Queen’s (first test, calibration of sensors, installation procedure development, operation procedure development ) | Oct 2018
Cryostat pieces cleaning | Oct 2018
Cryostat pieces shipped to SNOLAB | till Dec 2018
Calibration system production | till Nov 2018
Preparation for Installation Review ph 2B (cryostat, suspension system, calibration system, DAQ + electronics) | till Jan 2019
Preparation for Operational Readiness Review | till March 2019
Cleanroom operation test (air quality test, operational test, working procedure improvement) | till June 2019
Radio-assay of materials used for CUTE facility production | during 2018-2019
Completion of CUTE background model | till June 2019

## Slide 14: CUTE tasks till Operation with SuperCDMS detectors (2)
Step | Time required
Calibration system production | till Nov 2018
133Ba gamma source encapsulation (encapsulation and assay) | till Nov 2018
Production of transportation system for cryostat | till Jan 2019
Get Light, Electrical and Ethernet connections ready to use | depends on SNOLAB
Develop DAQ architecture for underground/surface CUTE lab | till end 2018
Test all procedures , its correction based on the results of first run with/without G115 | before June 2019
Evaluation of required manpower for run preparation, running | Before June 2019

## Slide 15: CUTE facility commissioning before SuperCDMS detector operation
Test of electronics components | DCRC boards VIB board Cables Pi-filters Power supply DAQ
Test of data handling system | Data quality procedure Transfer of raw data from DAQ to surface Access to surface/underground from outside
Calibration system test | Counting rate Data flow during calibration Remote functioning
Background evaluation | Background rate at ROI
Suspension system test | Running in different environmental conditions Effect to the detector
Coordination between CUTE-SuperCDMS teams | Develop the protocol of info/data exchange Responsibility sharing

## Slide 16: Payload for CUTE
Task | Time Period
Testing of CUTE facility at SNOLAB Without detectors With G115 detector/TES chips Revision/fixing problems/commissioning | March 2019 – June 2019
Si HV Pathfinder detector test | July 2019 – August 2019
First SuperCDMS tower (6 iZIP Ge detector) | September 2019 – May 2020
Commissioning
Operation

## Slide 17: Copper pieces cleanning

## Slide 18: Cleaning of cryostat parts
Magnetic shield
Copper cans

## Slide 19: Suspension system assembly
damper
supporting frame
added masses
(about 400 kg)
pulse tube
drywell
electronics box
turbopump

## Slide 20: Current status
supporting frame
added masses
(about 400 kg)
accessories

## Slide 21: Current status
added masses
(about 400 kg)
Alignment plates